Chapter 22: Evaluating standalone performance of CAD

# 22.1: Introduction

This chapter is focused on computer-aided detection (CAD) in screening mammography. On a mammogram there are four signs associated with breast cancer1. The two most common are masses (56%) and microcalcifications (29%). Less common are asymmetries (12%) and architectural distortions (4%). The majority (95%) of cases presenting as masses are invasive cancers, i.e., have spread from original location of the cancer and are associated with greater mortality, and the majority (68%), presenting as microcalcifications, are associated with ductal carcinoma in-situ (i.e., have not spread from the milk ducts). Mammography screening is a proven technique to reduce mortality from breast cancer2, annually about3 40,000 patients in the US, but mammographic screening, particularly detecting masses, is one of the most difficult tasks in radiology. There are a few reasons for this: the (fortunate) low incidence of breast cancer, about 5 per 1000, the lack of anatomic landmarks in the breast, the difficulty of imaging near the chest wall, and the camouflaging effect of dense (glandular) breast tissues. Typical sensitivity is about 80% with wide variability4,5, as much as 40%; a similar variability is observed for specificity, with experts achieving about 90% specificity. As a result, a woman's chance of a correct diagnosis depends on the radiologist who interprets her images.

Novel imaging treatments like breast computerized tomography6 and breast tomosynthesis7 have become realities. A major effort has been underway for almost three decades to develop computer aided detection (CAD) algorithms to analyze images for signs of cancer 8-10. CAD complements technological innovations because it can be applied to images from the newer modalities where performance, while better than conventional 2D images, is still far from perfect11. Almost all major imaging device manufacturers provide CAD as part of their imaging workstations. In the US the majority of digital mammograms are now analyzed by CAD12. In the UK CAD is not widely used, instead the radiologists depend on double reading.

A typical CAD algorithm consists of two stages: (i) a search stage, where image processing techniques are used to identify suspicious regions and (ii) a lesion-classification stage where features (quantitative descriptors) of the suspicious regions, such as size, contrast, shape, texture, etc.) are analyzed by a binary classifier (i.e., benign vs. malignant) to yield a malignancy index that an identified suspicious region is an actual cancer. If the index exceeds a pre-selected threshold, set by the manufacturer based on a compromise between sensitivity and specificity, the suspicious region in question is marked (or "annotated" or "cued"). After an initial unaided interpretation of the mammogram the radiologist reviews the annotations and reinterprets the mammogram: this is the current FDA approved 2nd reader mode for usage of CAD.

While there is consensus that CAD works for microcalcifications (regions containing small high contrast speck-like lesions), where it relieves the radiologist from the tedium of having to exhaustively search images with a magnifying glass, for masses (larger and relatively lower contrast lesions) its performance is controversial as evidenced by dueling editorials on the merits of CAD13,14 Two very large clinical studies15,16 (222,135 and 684,956 women, respectively) showed that CAD can actually have a detrimental effect on patient outcome13. These publications led to a flurry17-24 of rebuttals from the CAD community, mainly questioning the methodology used in these studies, one of which, relating to the use of clinical ratings for ROC studies was discussed in Section 4.10. A more recent study has confirmed the negative view of the efficacy of CAD25 and there has even been a call for ending Medicare reimbursement for CAD interpretations26.

Why the poor performance of CAD, even after 25 years of research? One reason has already been identified: excessive reliance on a poor FOM, namely FROC curve based measures. Another reason is lack of focus on the most obvious measure of the relative performance of CAD, namely a direct comparison of CAD to that of a group of experts. Clinical CAD systems sometimes only report the locations of suspicious regions, i.e., it may not provide ratings. However, a (continuous variable) malignancy index for every CAD-found suspicious region is available to the CAD algorithm designer27. Standalone performance, i.e., performance of CAD by itself, regarded as an algorithmic "reader", vs. a group of expert radiologists, is rarely measured. The author knows of only one study28 in breast CAD where standalone CAD performance was measured; other examples are to be found in CT colonography CAD and other modalities29-33.

The argument made for not measuring standalone performance is that since CAD will be used only as a 2nd reader, one need only measure performance of a radiologist without and with CAD. Indeed, a CAD expert has stated:34 “high stand-alone performance is neither a necessary nor a sufficient condition for CAD to be truly useful clinically.” Unfortunately, measuring performance with and without CAD has set a low bar for CAD to be considered useful. For example, CAD is not penalized for missing lesions as long as the radiologist finds them and CAD is not penalized for excessive FPs as long as the radiologist ignores them. There are also serious methodological problems, with excessive reliance on the low statistical power ROC paradigm and wide usage of the FROC curve as a measure of performance.

The purpose of this chapter is to describe a method to compare standalone performance of CAD to a group of radiologists interpreting the cases. The paradigm used to collect the data (e.g., ROC, FROC, LROC or ROI) is irrelevant – all that is needed is a scalar performance measure, or figure of merit (FOM), for the actual paradigm used. Before proceeding, in this chapter all references to CAD are to standalone CAD with designer-level data available. The author is grateful to Prof. Nico Karssemeijer for providing a dataset on standalone performance of CAD, without which this chapter would not be possible.

# 22.2: The Hupse-Karssemeijer et al. study

A study28 (henceforth referred to as the Hupse-Karssemeijer study) compared standalone performance of a CAD device to that of 9 radiologists interpreting the same cases (120 non-diseased and 80 with a single malignant mass per diseased case) using the location receiver operating characteristic (LROC) paradigm35-38, see Online Chapter 24 for summaries of all included datasets. In LROC, for each case the radiologist gives an overall rating for presence of disease, i.e., an ROC rating, and indicates the location of the most suspicious region. On non-diseased cases the rating is classified as a false positive (i.e., the location is ignored), but on a diseased case the rating is classified as a correct localization if the mark is sufficiently close to the lesion, and otherwise it is classified as an incorrect localization.

For a given threshold, the total number of correct localizations divided by the total number of diseased cases is an estimate of the probability of correct localization on a diseased case, denoted PCL. On non-diseased cases, the total number of false positives divided by the number of non-diseased cases is an estimate of the probability of a false positive, i.e., FPF. In the following equation, CL denotes correct localization, t denotes the truth status, denotes the chosen threshold and  denotes the (random) decision variable for case .

.

The plot of PCL (ordinate) vs. FPF defines the LROC plot. In the Hupse-Karssemeijer study "quasi-continuous" 0-100 integer ratings were used and the empirical LROCs were used to visually summarize performance.

Hupse-Karssemeijer used as FOMs the interpolated PCL at two values of FPF, namely 0.05 and 0.2. The two FOMs are denoted by and , respectively. Note that these figures of merit are point measures, as they depend on a single point on the LROC curve; in contrast, full area measures, such as AUC, depend on the entire curve (partial area measures are also possible, for example the area under the curve to the left of some pre-specified value of FPF). The t-test between the observed radiologist values and that of CAD was used to compute the (two-sided) p-value for rejecting the NH of equal performance. Hupse-Karssemeijer reported p-value = 0.17 for  and p-value < 0.001, with CAD being inferior, for . Table 22.1 summarizes relevant results from the Hupse-Karssemeijer publication (Table 1 ibid.). In the published table the ordinate of the LROC is labeled TPF, whereas in LROC notation38 it is customary to label it PCL.

The next section presents R code that duplicates the Hupse-Karssemeijer analysis.

Table 22.1: Hupse-Karssemeijer analysis with cases treated as a fixed factor. Note the different conclusions depending upon the choice of FPF value at which to evaluate the partial curve measure; <> denotes an average.

| FOM | <RAD> | CAD |  |  | p-value |
| --- | --- | --- | --- | --- | --- |
|  | 0.518 | 0.487 | 0.031 | (0.45,0.59) | 0.17 |
|  | 0.736 | 0.620 | 0.116 | (0.69,0.78) | <0.001 |

The paired t-test applied by Hupse-Karssemeijer treats the case sample as fixed. In other words, the analysis is not accounting for case-sampling variability but it is accounting for reader variability. As a historical / personal note, the author recalls a discussion (ca. 1986) with the late Prof. Charles E. Metz where Prof. Metz patiently explained to the author and Prof. Gary Barnes, his mentor at University of Alabama at Birmingham, that the (paired) t-test we were using in our Radiology publication39 was accounting for reader variability but not case variability. To paraphrase his comment: "I am comfortable with the claim that the observed significant difference generalizes to readers but it is specific to the one case sample that you studied." Since the 1986 study was a phantom study, the case sample happened to be the phantom with superposed simulated lesions. So the finding was actually specific to the one phantom that was used. Since a "population of phantoms" is not meaningful, therefore, even in retrospect, the paired t-test analysis used in the 1986 study, which disregarded case-variability, was appropriate. However, in the context of the Hupse-Karssemeijer study, it is necessary to account for case-sampling variability, i.e., patient-to-patient variability (to make the point explicit, patients are not phantoms; there is variability between them and a population of patients does exist).

## 22.2.1: Random-reader fixed-case analysis

The code and explanation to implement the Hupse-Karssemeijer analysis is in file mainAnalysisFixed.R, described in Online Appendix 22.A. The "fixed" in the file name emphasizes that case is regarded as a fixed factor in the analysis. At line 11 the variable named FOM is set to "PCL", which means the figure of merit is the interpolated PCL at the value of FPF specified at line 21, i.e., FPF = 0.05 on the first iteration of the for-loop starting at line 20. In other words the figure of merit for the first iteration is . Insert a break point at line 53 and Source the file to get Code Output 22.2.1.1 and Fig. 22.1.

### 22.2.1.1: Code Output

> debugSource(…)

Hupse-Karssemeijer analysis: random readers fixed cases:

FOM =  PCL

FPF =  0.05

Radiologists only

FomCad =  0.4625

AVG radiologist performance =  0.5063657 , 95%CI =  0.4368152 0.5759163

AVG diff. performance =  0.04386574 , 95%CI =  -0.02568478 0.1134163

t-statistic =  1.454404

df =  8

p-value =  0.1839158

...

Code Output 22.2.1.1 shows that FOM CAD, namely  is smaller than the average of the radiologists (0.463 < 0.506), but the p-value is 0.18, so the difference is not significant at . This result is mirrored by the fact that the 95% CI for the radiologists (0.44, 0.58) includes the figure of merit of CAD (0.463). The output lists the value of the t-statistic (1.45) and the degrees of freedom (8), which is one less than the number of radiologists.

Use the first command in the code snippet below to examine thetajc, the figure of merit array (j is the reader index and c is the case sample index, see Eqn.  below). The first reader is standalone CAD and the subsequent 9 readers correspond to the radiologists. The second command shows the difference between CAD and the average of the radiologists (the negative index removes CAD), with CAD being worse (-0.0439).

### 22.2.1.2: Code snippet

Browse[2]> thetajc

[1] 0.4625000 0.4250000 0.5156250 0.6750000 0.5291667 0.4479167 0.5979167 0.5375000 0.4250000 0.4041667

Browse[2]> mean(thetajc[-1]-thetajc[1])

[1] 0.04386574

The radiologists are slightly (0.044) better as a group than CAD. The question one wishes to answer is: is the difference statistically significant? Related questions are what is the p-value and what is the confidence interval for the FOM difference?

Figure 22.1 shows empirical LROC plots for the 9 radiologists and standalone CAD (heavy black line). Unlike ROC curves, the LROC curves flatten out at a value of PCL < 1. This behavior should be fairly obvious because even as the threshold is lowered, lesions that are not found by the search mechanism will not be correctly localized, especially since the reader is only allowed one mark. Exit debug mode, remove any breakpoints and source the code. This time all iterations in the for-loop are executed, where the iterations correspond to different FPF values at which to evaluate PCL. The code output was used to populate the rows labeled "fixed" in Table 22.2. These correspond to the values of FPF defined in the array at line 12, namely 0.05, 0.2, 0.5, 1. The table is split into four rows corresponding to different choices of FPF and each row is further split into two rows, corresponding to treating cases as fixed and random, respectively. All rows labeled "fixed" were obtained by sourcing mainAnalysisFixed.R, while rows labeled "random" were obtained by sourcing mainAnalysisRandom.R, to be described below. Instances where direct comparisons with the Hupse-Karssemeijer analysis are possible are indicated with asterisks.

At the two instances where direct comparison with the Hupse-Karssemeijer analyses is possible, i.e., fixed case analysis with PCL @ FPF = 0.05 and PCL @ 0.2, respectively, the agreement between 95% CIs and p-values is quite good. For example, @ FPF = 0.05, the CIs are (0.45, 0.59) in Table 22.1 vs. (0.44, 0.58) in Table 22.2. The minor differences in FOMs (Table 22.2 yields consistently smaller values than those in the Hupse-Karssemeijer publication) could be due to differences between the dataset actually used for the reported study and that received by us and / or methods for interpolating the data.

It should be fairly obvious from these plots that the radiologists, as a group, outperformed CAD. The difference is especially striking for the third radiologist. However, this straightforward conclusion can be obfuscated by focusing on the very low end of the plots, e.g., FPF = 0.05, where the plots are very steep, resulting in a not-significant p-value (i.e., 0.18). Some CAD developers use a point measure at the low end of FPF in the mistaken belief that this is the "clinically relevant" portion of the curve; this issue was discussed at length in §17.10.1. Appropriate analysis, using whole area measures of performance, is described in §22.5.

The output of the analysis for the FPF = 0.2 loop is shown below.

Fig. 22.1: Empirical LROC plots; the thick black plot applies to standalone CAD while the rest apply to the nine radiologists. This figure, which shows the entire operating characteristic, leads to the inescapable conclusion that the radiologists as a group outperformed CAD. Whether the difference is significant is determined by the analysis. This figure was generated by mainAnalysisFixed.R.

### 22.2.1.3: Code Output

...

FOM =  PCL

FPF =  0.2

FomCad =  0.5917

AVG radiologist performance =  0.7111

95%CI =  0.6702 0.7519

AVG diff. performance =  0.1194

95%CI =  0.07858 0.1603

t-statistic =  6.742

df =  8

p-value =  0.0001463

...

As expected from Fig. 22.1, at this FPF the difference (between the average of the radiologists and CAD) is significant, p-value = 0.00015. The average difference in FOMs is much larger, 0.119 vs. 0.044 in the previous example, and the 95% confidence interval for the radiologists (0.67, 0.75) now does not include the FOM for CAD (0.59). However, as emphasized previously, claiming that CAD is inferior based on a point FOM would be incorrect as the choice of FPF is arbitrary.

The Hupse-Karssemeijer analysis treats the case sample as fixed. The following section extends the analysis to random cases.

# 22.3: Extending the analysis to random cases

Standalone CAD is an algorithmic "reader", not a different treatment. Therefore a method for analyzing readers in a single treatment 40,41 such a described in §10.7, is appropriate, where the figure of merit for the radiologists was modeled as (j is the reader index, excluding CAD):

.

is the figure of merit for radiologist j (j = 1, 2, ..., J) interpreting case sample ;  is the random effect of radiologist j and is the error term described in Chapter 10 and below, Eqn. .

The formulae presented in §10.7, allow one to test the NH:  where  is a pre-specified constant. One could set  equal to the standalone performance of CAD, but that would ignore the fact that the performance of CAD is itself a random variable, not a constant, whose case-sampling variability needs to be accounted for.

Instead, the following model is used for the figure of merit fluctuations of the radiologists and CAD:

.

is the CAD figure of merit estimate for case sample , i.e.,  j = 0 is used to denote the CAD "reader";  is the average figure of merit increment of the radiologists over CAD. Subtract the CAD figure of merit from the radiologist figure of merit for the same case sample, and define this as the difference (radiologist minus CAD) figure of merit , i.e.,

.

Then Eqn.  can be written as:

.	.

Eqn.  is homologous (i.e., identical in form) to Eqn.  with the subtle difference that the figure of merit on the left hand side of Eqn.  is a difference FOM, between the radiologist's FOM and that of CAD. Eqn.  describes a model for J difference radiologists interpreting a common case set , each of whose performances is a difference from that of CAD (the difference is positive if the radiologist is better). Under the NH the expected difference is zero: . The method described in §10.7 for single treatment multiple reader analysis is now directly applicable to the model described by Eqn. . One assumes that the error term is distributed as:

.

The J x J covariance matrix  is defined by two parameters, and , as follows:

.

The terms and  can be estimated using resampling methods, e.g., the jackknife, see Chapter 07 and Chapter 10. Denote the difference FOM with case k removed by  (the index in parenthesis  denotes deleted case k; i.e., it is the FOM with case k removed from case set . For notational conciseness, and since one is restricted to c = 1, the case set index is henceforth suppressed. The covariance matrix is estimated using (the variance inflation factor is shown explicitly):

.

The final estimates of and  are averaged (indicated by the angular brackets below) over all pairings of radiologists satisfying the relevant equalities/inequalities shown just below the closing angular bracket:

.

Under the null hypothesis  the observed value of the statistic  defined below follows the t-distribution with  degrees of freedom, where is defined in Eqn.  (compare to Equation 10.72):

.

Here is the observed effect size, i.e., the average, over radiologists, of , Eqn. , and  is defined by (compare to Equation 10.68):

.

The degrees of freedom of the t-statistic is defined by (compare to Eqn. 10.72):

.

The 2-sided p-value for rejecting the NH:  is:

.	.

A  percent confidence interval for the observed effect size is given by40 (compare to Eqn. 10.74):

.

Here  is the upper  quantile of , i.e., the value such that  of the specified distribution lies above it.

Open the file mainAnalysisRandomBrief.R, a listing of which follows (Brief in the file name signifies that this file is a brief version, a more general version handles other FOMs).  Random in the file name implies that cases are treated as random, as in the analysis presented in Eqn.  – Eqn. . In the brief version the FOM is set to AUC, the empirical area under the ROC curve; FPF is then irrelevant, but it is needed to use other FOMs. In case one wonders how to get ROC data from LROC, one uses the maximum rating, which could be a correct or an incorrect localization event (see the pmax() function in line 20-25 of the following code listing; this stands for "parallel maximum" over the correct and incorrect localization arrays, each of length ).

### 22.3.1: Code Listing

# MainAnalysisRandomBrief.R

rm(list = ls())

library(RJafroc);library(ggplot2)

source("DiffFomAnal2007Hillis53.R");source("Wilcoxon.R")

alpha <- 0.05

FOM <- "AUC"

FPF <- 0.0

cat("FOM = ", FOM, "\n")

if (FOM == "PCL") cat("FPF = ", FPF, "\n")

cat("Random-reader random-case analysis")

cat("\nof Hupse Karssemeijer radiologist data:\n")

retNico <- DfReadLrocDataFile()

zjk1 <- retNico$NL[1,,,1]

zjk2Cl <- retNico$LLCl[1,,,1]

zjk2Il <- retNico$LLIl[1,,,1]

zjk2 <- pmax(zjk2Cl,zjk2Il)

K1 <- length(zjk1[1,])

K2 <- length(zjk2Cl[1,]);K <- c(K1,K2)

ret_nh <- DiffFomAnal2007Hillis53 (

alpha,

zjk1,

zjk2,

FOM,

FPF)

cat(str(ret_nh))

DiffFomAnal2007Hillis53() implements the analysis described in Equations 22.3 through 22.14 and gives due credit to Section 5.3 of a Hillis publication41. Source the file to get the following output.

### 22.3.2: Code Output

> source('~/book2/05 D Advanced Topics/D22 CAD vs. radiologists/software/mainAnalysisRandomBrief.R')

FOM =  AUC

Random-reader random-case analysis of Hupse Karssemeijer data:

List of 9

$ Var    : num 0.0014

$ Cov2   : num 0.000924

$ PsiMean: num 0.0317

$ CI     : num [1:2(1d)] -0.031 0.0945

$ reject : num 0

$ ddfH   : num 878

$ Tstat  : num 0.993

$ p_val  : num 0.321

$ thetajc: num [1:10(1d)] 0.817 0.842 0.841 0.9 0.838 ...

The lines starting with $Var and $Cov2 show the two elements of the covariance matrix:  = 0.0014 and = 0.000924. Covariance is smaller than variance, which should make sense, since covariance equals the variance times a correlation, and the latter is less than unity. The next line lists the mean value of the difference figure of merit, PsiMean which stands for  = 0.0317 (the difference in empirical AUCs of the radiologists minus that of CAD, averaged over radiologists). The next line lists the 95% confidence interval for , which is [-0.031, 0.0945]. Since this includes zero, the NH, that average differential performance is zero, is not rejected, emphasized by the next line reject = 0. The next line lists the degrees of freedom of the t-test, = 878; the  statistic is 0.993, the p-value is 0.321 and the last line prints thetatjc (): the 1st value corresponds to CAD; the 2nd value corresponds to the 1st radiologist, etc.

### 22.3.3: Extension to more FOMs

The file mainAnalysisRandom.R, listed in Online Appendix 22.B.1, extends mainAnalysisRandomBrief.R to other FOMs. Ensure that FOM is set to "PCL" at line 10. Source the code to get the random-case results summarized in Table 22.2. Because it is accounting for an additional source of variability, each of the lines labeled random yields a larger p-value and a wider confidence interval than the corresponding line labeled fixed.

Table 22.2: Comparison of fixed case and random case analyses. All rows labeled "fixed" were obtained by sourcing mainAnalysisFixed.R, while rows labeled "random" were obtained by sourcing mainAnalysisRandom.R. The figure of merit (FOM) = PCL @ specified FPF. An asterisk indicates instances where direct comparison with Hupse-Karssemeijer analyses is possible. Because it is accounting for an additional source of variability, each of the lines labeled random yields a larger p-value and a wider confidence interval than the corresponding line labeled fixed.

| FPF | Cases | CAD | <RAD> |  |  |  | p-value |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0.05 | fixed | 0.463* | 0.506* | (0.44,0.58)* | 0.044 | (-0.026,0.11) | 0.18* |
|  | random |  |  | (0.36,0.65) |  | (-0.100 0.188) | 0.55 |
| 0.2 | fixed | 0.592* | 0.711* | (0.67,0.75)* | 0.119 | (0.079, 0.16) | 0.0001* |
|  | random |  |  | (0.60,0.83) |  | (0.004, 0.235) | 0.042 |
| 0.5 | fixed | 0.675 | 0.775 | (0.73,0.82) | 0.100 | (0.059, 0.141) | 0.0005 |
|  | random |  |  | (0.67,0.88) |  | (-0.004, 0.204) | 0.059 |
| 1 | fixed | 0.675 | 0.783 | (0.74,0.83) | 0.108 | (0.065, 0.152) | 0.0004 |
|  | random |  |  | (0.68,0.89) |  | (0.005, 0.212) | 0.041 |

# 22.4: Ambiguity in interpreting a point-based FOM

Both analyses, i.e., fixed-case or random-case, in Table 22.2 declare not significant difference at FPF = 0.05 and significant difference at FPF = 0.2. The difference in conclusions between FPF = 0.05 and FPF = 0.2 illustrates the ambiguity when using point-based measures, which do not use all of the data. For example, at FPF = 0.05, most of the data contributing to the plots to the right of this operating point, Fig. 22.1, are not used in the analysis. An even more important reason is the steepness of the curves at this low FPF, which introduces greater variability in the interpolated estimates.

The variable results of point measures, such as used in Table 22.1 and Table 22.2, invites "gaming" the system; i.e., the researcher is under temptation to report performance at a FPF value where the algorithm "looks good". Using FPF = 0.05 the designer could claim (incorrectly as it turns out, because inability to reject the NH does not mean the NH is true) their method is no worse than experts. Another investigator using FPF = 0.2 could claim that the algorithm is worse than radiologists (a correct conclusion as it turns out, but incorrect method, because the p -value depends on the choice of FPF, which is arbitrary).

# 22.5: Results using full-area measures

Table 22.3 presents results for choices of FOM, each of which is a full area measure. A full-area measure removes the arbitrariness inherent in any partial area or point-based measure. Shown are results for the empirical area under the LROC plot (FOM = "ALroc" at line 10 in mainAnalysisRandom.R) and the empirical area under the ROC plot (FOM = "AUC").

Table 22.3: This table illustrates the result of using area empirical FOM measures. Two FOMs are used, as indicated in the first column.

| FOM | Cases | CAD | <RAD> |  |  |  | p-value |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | fixed | 0.628 | 0.734 | (0.69,0.77) | 0.105 | (0.065, 0.146) | 0.0003** |
|  | random |  |  | (0.64,0.83) |  | (0.008, 0.203) | 0.0349 |
|  | fixed | 0.817 | 0.849 | (0.83,0.87) | 0.0317 | (0.009, 0.055) | 0.0124** |
|  | random |  |  | (0.79,0.91) |  | (-0.031, 0.094) | 0.321 |

Legend: **Significant at alpha corrected for multiple testing, i.e., test alpha = 0.025.

Before getting into a discussion of these results, it is appropriate to note that a correction for multiple testing is needed when applying multiple analyses to the same dataset. One way of accomplishing this is by setting the alpha for multiple testing at the nominal alpha divided by the number of tests. In other words, to maintain the overall probability of a Type I error at 0.05, one needs to adopt the multiple testing alpha as 0.05/2, i.e., 0.025. This is often referred to as the Bonferroni correction42,43. The Bonferroni correction is conservative – it is accurate when the data are independent – but in the presence of correlations it would lead to inability to reject the NH more often than is appropriate (i.e., loss of statistical power). It has the advantage of being the simplest method for accounting for multiple comparisons.

Using the total area under the empirical LROC (1st results row in Table 22.3) yields a not-significant difference for random cases (p = 0.0349 > 0.025; the use of 0.025 for alpha is explained below), but a significant difference for fixed cases (p = 0.0003 < 0.025), with CAD being inferior, consistent with the visual impression of the plots in Fig. 22.1. Similarly, the corresponding analysis using the empirical area under the ROC yields a significant difference only if cases are regarded as fixed (p=0.0124 < 0.025). Both full area measures, LROC and ROC, are in agreement, yielding a significant difference only if case is regarded as a fixed factor. However, the p-value for LROC is much smaller (0.0003 << 0.0124). This is consistent with greater statistical power using LROC, which takes location into account, as compared to ROC. The increased statistical power has been noted by Swensson38,44,45.

If prior to conducting the analysis one had committed to either FOM, not both, then one would be justified in not applying the Bonferroni correction. In the current situation, based on the expectation of higher power, the author would have chosen the full area under the LROC, in which case the difference is significant even when case is regarded as a random factor (p=0.0349 < 0.05). Of course, having committed, scientific integrity requires that one does not change course after the study is completed, just because the results look better using an alternate FOM. If one does change course, then the Bonferroni correction is needed. It is there to keep one honest.

# 22.6: Discussion / Summary

This chapter has described an extension of the Hupse-Karssemeijer analysis that accounts for case variability. The method extends the single treatment analysis of Chapter 10 to a situation where one of the "readers" is a special reader, and the desire is to compare performance of this reader against the average of the rest of the readers. The method was used to analyze, using different figures of merit, the LROC data set kindly provided by Prof. Karssemeijer. It is shown that point-based measures of performance, such as PCL @ specified FPF, lead to different results, depending on the choice of FPF. Similar conclusions apply to partial area based FOMs. The only way of avoiding this is to use full-area based measures, such as the area under the ROC or the LROC. An added advantage of doing so is the increased statistical power due to the fact that one is using all the data; point-based measures or partial-area measures do not use all the data. The currently encouraged focus on the low FPF end of the ROC curve is misleading, as one is trying to measure performance on the initial steep area of the curve, where there is little data, and most of the data that exists, comes from highly visible lesions, which both CAD and the radiologists are expected to find easily, leading to further loss of statistical power. Using a full-area measure takes all lesions in the dataset into account, and the findings extrapolate to the population of cases, not just to the subset of cases with highly visible lesions.

While demonstrated using ROC and LROC data, the method is applicable to any scalar figure of merit. It is recommended that future CAD studies use FROC methodology in conjunction with the empirical wAFROC-AUC figure of merit. The advantage of using the FROC paradigm lies specifically in higher statistical power and the diagnostics yielded by RSM-based ROC curve fitting; specifically the RSM parameters yield information on what is limiting CAD performance – i.e., search or classification performance. Based on the author's experience current CAD algorithms are much worse than experts in search performance (e.g., CAD generates 30 times the number of NLs per 4-view case than experts), but CAD's classification performance might actually be better (explained by the fact that it has access to the digital pixel values and the ability to implement sophisticated feature-detection and lesion-classification methods).

# 22.7: References

1.	Shetty MK. Breast Cancer Screening and Diagnosis: A Synopsis. Springer; 2014.

2.	Kerlikowske K, Grady D, Rubin SM, Sandrock C, Ernster VL. Efficacy of screening mammography: a meta-analysis. Jama. 1995;273(2):149-154.

3.	ACS. American Cancer Society: Cancer Facts & Figures 2015 Atlanta: American Cancer Society; 2015.

4.	Beam CA, Layde PM, Sullivan DC. Variability in the interpretation of screening mammograms by US radiologists. Findings from a national sample. Archives of Internal Medicine. 1996;156(2):209-213.

5.	Elmore JG, Jackson SL, Abraham L, Miglioretti DL, Carney PA, Geller BM, Yankaskas BC, Kerlikowske K, Onega T, Rosenberg RD, Sickles EA, Buist DSM. Variability in Interpretive Performance at Screening Mammography and Radiologists’ Characteristics Associated with Accuracy. Radiology. 2009;253(3):641-651.

6.	Boone JM, Kwan ALC, Yang K, Burkett GW, Lindfors KK, Nelson TR. Computed Tomography for Imaging the Breast. J Mammary Gland Biol Neoplasia. 2006;11:103–111.

7.	Niklason LT, Christian BT, Niklason LE, Kopans DB, Castleberry DE, Opsahl-Ong BH, Landberg CE, Slanetz PJ, Giardino AA, Moore R, Albagli D, DeJule MC, Fitzgerald PF, Fobare DF, Giambattista BW, Kwasnick RF, Liu J, Lubowski SJ, Possin GE, Richotte JF, Wei CY, Wirth RF. Digital tomosynthesis in breast imaging. Radiology. 1997;205(2):399-406.

8.	Giger ML, Chan H-P, Boone J. Anniversary Paper: History and status of CAD and quantitative image analysis: The role of Medical Physics and AAPM. Medical Physics. 2008;35(12):5799-5820.

9.	Nishikawa RM. Current status and future directions of computer-aided diagnosis in mammography. Computerized Medical Imaging and Graphics. 2007;31(4–5):224-235.

10.	Li Q, Nishikawa RM, eds. Computer-Aided Detection and Diagnosis in Medical Imaging. Taylor & Francis; 2012. Hendee WR, ed. Imaging in Medical Diagnosis and Therapy.

11.	Chan H-P, Wei J, Sahiner B, Rafferty EA, Wu T, Roubidoux MA, Moore RH, Kopans DB, Hadjiiski LM, Helvie MA. Computer-aided Detection System for Breast Masses on Digital Tomosynthesis Mammograms: Preliminary Experience. Radiology. 2005;237(3):1075-1080.

12.	Rao VM, Levin DC, Parker L, Cavanaugh B, Frangos AJ, Sunshine JH. How widely is computer-aided detection used in screening and diagnostic mammography? J Am Coll Radiol. 2010;7:802–805.

13.	Philpotts LE. Can Computer-aided Detection Be Detrimental to Mammographic Interpretation? Radiology. 2009;253(1):17-22.

14.	Birdwell RL. The Preponderance of Evidence Supports Computer-aided Detection for Screening Mammography1. Radiology. 2009;253(1):9-16.

15.	Fenton JJ, Taplin SH, Carney PA, Abraham L, Sickles EA, D'Orsi C, Berns EA, Cutter G, Hendrick RE, Barlow WE, Elmore JG. Influence of Computer-Aided Detection on Performance of Screening Mammography. N Engl J Med. 2007;356(14):1399-1409.

16.	Fenton JJ, Abraham L, Taplin SH, Geller BM, Carney PA, D’Orsi C, Elmore JG, Barlow WE, Consortium ftBCS. Effectiveness of Computer-Aided Detection in Community Mammography Practice. Journal of the National Cancer Institute. 2011.

17.	Ciatto S, Houssami N. Computer-Aided Screening Mammography. New England Journal of Medicine. 2007;357(1):83-85.

18.	Feig SA, Birdwell RL, Linver MN. Computer-Aided Screening Mammography. New England Journal of Medicine. 2007;357(1):83-85.

19.	Fenton J, J, Barlow WE, Elmore JG. Computer-Aided Screening Mammography. New England Journal of Medicine. 2007;357(1):83-85.

20.	Gur D. Computer-Aided Screening Mammography. New England Journal of Medicine. 2007;357(1):83-85.

21.	Nishikawa RM, Schmidt RA, Metz CE. Computer-Aided Screening Mammography. New England Journal of Medicine. 2007;357(1):83-85.

22.	Ruiz JF. Computer-Aided Screening Mammography. New England Journal of Medicine. 2007;357(1):83-85.

23.	Berry DA. Computer-Assisted Detection and Screening Mammography: Where’s the Beef? Journal of the National Cancer Institute. 2011.

24.	Nishikawa RM, Giger ML, Jiang Y, Metz CE. Re: Effectiveness of Computer-Aided Detection in Community Mammography Practice. Journal of the National Cancer Institute. 2012;104(1):77.

25.	Lehman CD, Wellman RD, Buist DS, Kerlikowske K, Tosteson AN, Miglioretti DL. Diagnostic accuracy of digital screening mammography with and without computer-aided detection. JAMA internal medicine. 2015:1-10.

26.	Fenton JJ. Is It Time to Stop Paying for Computer-Aided Mammography? JAMA Intern Med. 2015.

27.	Edwards DC, Kupinski MA, Metz CE, Nishikawa RM. Maximum likelihood fitting of FROC curves under an initial-detection-and-candidate-analysis model. Med Phys. 2002;29(12):2861-2870.

28.	Hupse R, Samulski M, Lobbes M, Heeten A, Imhof-Tas M, Beijerinck D, Pijnappel R, Boetes C, Karssemeijer N. Standalone computer-aided detection compared to radiologists’ performance for the detection of mammographic masses. Eur Radiol. 2013;23(1):93-100.

29.	Hein PA, Krug LD, Romano VC, Kandel S, Hamm B, Rogalla P. Computer-aided detection in computed tomography colonography with full fecal tagging: comparison of standalone performance of 3 automated polyp detection systems. Canadian Association of Radiologists Journal. 2010;61(2):102-108.

30.	Summers RM, Handwerker LR, Pickhardt PJ, Van Uitert RL, Deshpande KK, Yeshwant S, Yao J, Franaszek M. Performance of a previously validated CT colonography computer-aided detection system in a new patient population. American Journal of Roentgenology. 2008;191(1):168-174.

31.	Taylor SA, Halligan S, Burling D, Roddie ME, Honeyfield L, McQuillan J, Amin H, Dehmeshki J. Computer-assisted reader software versus expert reviewers for polyp detection on CT colonography. American Journal of Roentgenology. 2006;186(3):696-702.

32.	De Boo DW, Uffmann M, Weber M, Bipat S, Boorsma EF, Scheerder MJ, Freling NJ, Schaefer-Prokop CM. Computer-aided detection of small pulmonary nodules in chest radiographs: an observer study. Academic radiology. 2011;18(12):1507-1514.

33.	Tan T, Platel B, Huisman H, Sánchez C, Mus R, Karssemeijer N. Computer-aided lesion diagnosis in automated 3-D breast ultrasound using coronal spiculation. Medical Imaging, IEEE Transactions on. 2012;31(5):1034-1042.

34.	Nishikawa RM, & Pesce, L. L. Fundamental limitations in developing computer-aided detection for mammography. Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment. 2011(648):S251–S254.

35.	Metz CE, Starr SJ, Lusted LB. Observer performance in detecting multiple radiographic signals. Prediction and analysis using a generalized ROC approach. Radiology. 1976;121(2):337-347.

36.	Starr SJ, Metz CE, Lusted LB, Goodenough DJ. Visual detection and localization of radiographic images. Radiology. 1975;116:533-538.

37.	Swensson RG, Judy PF. Detection of noisy visual targets: Models for the effects of spatial uncertainty and signal-to-noise ratio. Perception & Psychophyics. 1981;29(6):521-534.

38.	Swensson RG. Unified measurement of observer performance in detecting and localizing target objects on images. Med. Phys. 1996;23(10):1709 -1725.

39.	Chakraborty DP, Breatnach ES, Yester MV, Soto B, Barnes GT, Fraser RG. Digital and Conventional Chest Imaging: A Modified ROC Study of Observer Performance Using Simulated Nodules. Radiology. 1986;158:35-39.

40.	Hillis SL. A marginal‐mean ANOVA approach for analyzing multireader multicase radiological imaging data. Statistics in medicine. 2014;33(2):330-360.

41.	Hillis SL. A comparison of denominator degrees of freedom methods for multiple observer ROC studies. Statistics in Medicine. 2007;26:596-619.

42.	Perneger TV. What’s wrong with Bonferroni adjustments. BMJ: British Medical Journal. 1998;316(7139):1236.

43.	Bland JM, Altman DG. Multiple significance tests: the Bonferroni method. Bmj. 1995;310(6973):170.

44.	Swensson RG, King JL, Good WF, Gur D. Observer variation and the performance accuracy gained by averaging ratings of abnormality. Medical Physics. 2000;27(8):1920-1933.

45.	Swensson RG. Using Localization Data from Image Interpretations to Improve Estimates of Performance Accuracy. Medical Decision Making. 2000;20(2):170-184.